Directions

- Initialize Asylum software, take 512x512 preliminary image

- Set lockin to ARC: Master Panel -> Tune -> Adv -> Which Lockin -> ARC

- AFM controls -> Litho Panel -> Bitmap -> Check “Bitmap Is Bias”

- Copy and paste procedure into procedure window (accessed with control+M)

- Open command window (Ctrl+J)

- SetDataFolder root:Packages:MFP3D:XPT:Cypher

- Click on “macros” from the toolbar, then “nanoRASP panel”

- Edit nRASP variables (see explanations below)

- Click “Load Excel Data”, navigate to the excel sheet you want

- Make sure the first tab of the excel sheet is the one you want to load

- In excel, a value of 0 means no dig, and a value of 255 means max dig

- Also, make sure target pattern is flipped vertically (it will end up flipping back when the image is taken)

- display; appendimage trgt_scaled

- Make sure img_num is the number of the image you want to compare to (usually the last image taken)

- Call getForce() from command window, make sure it looks good

- display; appendImage lith_force

- Check img_num again

- Programming -> User Callbacks Panel

- Check “enable callbacks” and “imageLastScan”

- Next to ImageLastScan, write InitCustomScan()

- Make sure Data Browser is closed

- Call InitCustomScan() from command window

Troubleshooting

Common bugs that come up:

- Once you have started litho, before doing any other scans, make sure you close the data browser

- Make sure img_num is the right value (i.e. refers to a file that exists). If it isn’t, you’ll have to close a few debugging/troubleshooting windows, decrement img_num, and try again

- If real time images aren't changing, but it is scanning, then the data browser was open when you started. Click on "igor abort" on bottom left (might be hidden by windows toolbar), then click "stop" to end scan. Close data browser and try again

nRASP Panel

- vthreshold: threshold voltage, where AFM actually starts digging

- vsp: setpoint voltage, where no digging occurs (just scanning)

- vmax: max voltage to be applied

- DIGPFR: Nanomaters dug per frame at vmax

- Kval: deflection coefficient

- Target depth: target contrast in image (nm)

- Border Width: Number of pixels surrounding digging area

- Border Width = (512 – image side length)/2

- Img_num: Suffix number of image to be compared to

- Total Images: number of images to take before terminating process

- Deflection channel: Channel for deflection data that will be used in comparison

- Height channel: Channel for height data that will be used in comparison

- Horizontal Shift: Indicates the number of pixels by which to offset the force application on the X axis. Positive value will shift where force is applied to the right (in relation to where the data is received)

- Vertical Shift: Indicates the number of pixels by which to offset the force application on the Y axis. Positive value will shift where force is applied up (in relation to where the data is received)